## export
model | G | D | B | zero_offset | calibration | temp_response | cosine_response | azimuth_response | tilt_response | fIAM | alias | type_default
GHI_A | 1 | 1 | 0 | 7.3485 | 2.9816 | 0.04 | 10 | 0 | 0.5 |  | (secondary[\W_]?standard)¦S¦A¦(high[\W_]?quality)
GHI_B | 1 | 1 | 0 | 16.3095 | 5.6789 | 0.08 | 20 | 0 | 2 |  | ((1¦1st¦I¦first)[\W_]?class)¦B¦(good[\W_]?quality) | GHI, USW
GHI_C | 1 | 1 | 0 | 32.619 | 11.0454 | 0.16 | 30 | 0 | 5 |  | ((2¦2nd¦II¦second)[\W_]?class)¦C¦(moderate[\W_]?quality) | GTI
CMP22 | 1 | 1 | 0 | 3.3166 | 2.0712 | 0.01 | 5 | 0 | 0.2 |  | cmp22
CMP21 | 1 | 1 | 0 | 7.3485 | 2.0712 | 0.02 | 10 | 0 | 0.2 |  | cmp21
CMP11 | 1 | 1 | 0 | 7.3485 | 2.0712 | 0.02 | 10 | 0 | 0.2 |  | cm(p¦a)11
CMP6 | 1 | 1 | 0 | 15.5563 | 2.4495 | 0.08 | 20 | 0 | 1 |  | cm(p¦a)6
MS80 | 1 | 1 | 0 | 1.99 | 2.0298 | 0.02 | 10 | 0 | 0.2 |  | ms80
PSP | 1 | 1 | 0 | 1 | 3.3541 | 0.02 | 30 | 2 | 0 | @(z) 1.036*cosd(z).^0.073 | psp
R | 1 | 1 | 0 | 1 | 4.6904 | 0.1 | 60 | 0 | 0 |  | (ref¦cal).*cell¦R
RSI | 1 | 1 | * | 1 | 3.3106 | 0.012 | 40 | 2.6 | 0 |  | rsi
LI200 | 1 | 1 | 0 | 1 | 6.245 | 0.02 | 40 | 2.6 | 0 | @LI200_IAM | (LI[\W_]?200)¦(photo.+diode)¦P
BNI_A | 0 | 0 | 1 | 2.1237 | 1.5969 | 0.02 | 0 | 0 | 0.2 |  | (secondary[\W_]?standard)¦S¦A¦(high[\W_]?quality)
BNI_B | 0 | 0 | 1 | 4.1533 | 1.9365 | 0.04 | 0 | 0 | 0.5 |  | ((1¦1st¦I¦first)[\W_]?class)¦B¦(good[\W_]?quality) | BNI
CM22 | 1 | 1 | 0 | 0.5 | 2.3524 | 0.01 | 0 | 1 | 0 |  | cm22
CM6b | 1 | 1 | 0 | 0.5 | 2.6567 | 0.04 | 0 | 2.2361 | 0 |  | cm6b
CM3 | 1 | 1 | 0 | 0.5 | 6.4226 | 0.12 | 0 | 2.2361 | 0 |  | cm3
TSP800 | 1 | 1 | 0 | 0.5 | 1.8975 | 0.02 | 0 | 2.2361 | 0 |  | tsp[\W_]?800
TSP1 | 1 | 1 | 0 | 0.5 | 3.2172 | 0.02 | 0 | 2.2361 | 0 |  | tsp[\W_]?1
SPN1 | 1 | 1 | * | 0.5 | 9.5974 | 0.02 | 0 | 2.2361 | 0 |  | spn1 | DHI
SPLite | 1 | 1 | 0 | 0.5 | 5.8732 | 0.003 | 0 | 2.2361 | 0 |  | SP[\W_]?lite
SP110 | 1 | 1 | 0 | 0.5 | 5.707 | 0.02 | 0 | 2.2361 | 0 |  | SP[\W_]?110
CH1 | 0 | 0 | 1 | 0.5 | 1.7659 | 0.02 | 0 | 2.2361 | 0 |  | ch1
NIP2 | 0 | 0 | 1 | 0.5 | 1.9996 | 0.02 | 0 | 1.118 | 0 |  | (NIP[\W#_]?2¦NIP)
NIP1 | 0 | 0 | 1 | 0.5 | 1.9577 | 0.02 | 0 | 0.5 | 0 |  | NIP[\W#_]?1
LI201 | 1 | 1 | 0 | 2 | 8.8396 | 0.02 | 0 | 0.5 | 0 |  | LI[\W_]?201
Schenk | 1 | 1 | 0 | 1 | 2.2913 | 0.02 | 30 | 3 | 2 |  | (shenk¦schenk)
CHP1 | 0 | 0 | 1 | 1.118 | 0.96437 | 0.01 | 0 | 0.5 | 0 |  | chp1
CM11 | 1 | 1 | 0 | 14.4568 | 2.9343 | 0.06 | 50 | 0 | 0.25 |  | cm11

## list
[6 empty] | BNI | 800 |  | zenith | 45
[6 empty] | DHI | 200 |  | temp | 40
[6 empty] | GHI | 765.6854 |  | ventilated | 0
[9 empty] | tilted | 0
[16 empty] | Zero offset |  |  |  | Calibration |  |  |  |  | Directional
|  | G | D | B | zero_offset | calibration | temp_response | cosine_response | azimuth_response | tilt_response | fIAM | Model alias (regexp) | Uncertainty GHI [W/m²] | Uncertainty DHI [W/m²] | Uncertainty DNI [W/m²] | Due to thermal radiation | Due to temperature | Resolution & Logger | Other | Tracing to WRR | Stability | Non-linearity | Spectral response | Other | Azimuth & other | Leveling & cosine ~ tan(z) (error @ 80°, 1000 W/m²) | Temperature response | Tilt response | Sources, notes
1 | GHI_A | 1 | 1 | 0 | 7.3485 | 2.982% | 0.04% | 10 | 0% | 0.5% |  | (secondary[\W_]?standard)¦S¦A¦(high[\W_]?quality) | 24.6616 | 9.4636 |  | 7 | 2 | 1 |  | 2% | 0.8% | 0.5% | 2% |  | 0% | 10 | 2% | 0.5% | [1] "High Quality" ~ ISO Secondary Standard
2 | GHI_B | 1 | 1 | 0 | 16.3095 | 5.679% | 0.08% | 20 | 0% | 2% |  | ((1¦1st¦I¦first)[\W_]?class)¦B¦(good[\W_]?quality) | 47.8405 | 19.8746 |  | 15 | 4 | 5 |  | 2% | 1.5% | 1% | 5% |  | 0% | 20 | 4% | 2% | [2] "Good Quality" ~ ISO first class
3 | GHI_C | 1 | 1 | 0 | 32.619 | 11.05% | 0.16% | 30 | 0% | 5% |  | ((2¦2nd¦II¦second)[\W_]?class)¦C¦(moderate[\W_]?quality) | 92.2686 | 39.3954 |  | 30 | 8 | 10 |  | 2% | 3% | 3% | 10% |  | 0% | 30 | 8% | 5% | [1] "Moderate Quality" ~ ISO Second Class
4 | CMP22 | 1 | 1 | 0 | 3.3166 | 2.071% | 0.01% | 5 | 0% | 0.2% |  | cmp22 | 16.4548 | 5.3066 |  | 3 | 1 | 1 |  | 2% | 0.5% | 0.2% | 0% |  | 0% | 5 | 0.5% | 0.2% | [2]
5 | CMP21 | 1 | 1 | 0 | 7.3485 | 2.071% | 0.02% | 10 | 0% | 0.2% |  | cmp21 | 18.3985 | 8.4356 |  | 7 | 2 | 1 |  | 2% | 0.5% | 0.2% | 0% |  | 0% | 10 | 1% | 0.2% | [2]
6 | CMP11 | 1 | 1 | 0 | 7.3485 | 2.071% | 0.02% | 10 | 0% | 0.2% |  | cm(p¦a)11 | 18.3985 | 8.4356 |  | 7 | 2 | 1 |  | 2% | 0.5% | 0.2% | 0% |  | 0% | 10 | 1% | 0.2% | [2]
7 | CMP6 | 1 | 1 | 0 | 15.5563 | 2.449% | 0.08% | 20 | 0% | 1% |  | cm(p¦a)6 | 26.9396 | 16.3095 |  | 15 | 4 | 1 |  | 2% | 1% | 1% | 0% |  | 0% | 20 | 4% | 1% | [2]
8 | MS80 | 1 | 1 | 0 | 1.99 | 2.03% | 0.02% | 10 | 0% | 0.2% |  | ms80 | 16.6883 | 4.5211 |  | 1 | 1 | 1.4 |  | 2% | 0.2% | 0.2% | 0.2% |  | 0% | 10 | 1% | 0.2%
9 | PSP | 1 | 1 | 0 | 1 | 3.354% | 0.02% | 30 | 2% | 0% | @(z) 1.036*cosd(z).^0.073 | psp | 32.9471 | 6.7823 |  |  |  | 1 |  | 3% | 1% | 0.5% | 1% |  | 2% | 30 | 1% | 0% | [7], http://solardat.uoregon.edu/EppleyPSP.html, own data analysis
10 | R | 1 | 1 | 0 | 1 | 4.69% | 0.1% | 60 | 0% | 0% |  | (ref¦cal).*cell¦R | 49.7857 | 9.434 |  |  |  | 1 |  | 2% | 1% | 1% | 4% |  | 0% | 60 | 5% | 0% | Guesstimate by inspection of [3]
11 | RSI | 1 | 1 | * | 1 | 3.311% | 0.012% | 40 | 2.6% | 0% |  | rsi | 37.2531 | 6.6963 |  |  |  | 1 |  | 2% | 1.4% | 1% | 2% |  | 2.6% | 40 | 0.6% | 0% | [4], [5], [6], [7]
12 | LI200 | 1 | 1 | 0 | 1 | 6.245% | 0.02% | 40 | 2.6% | 0% | @LI200_IAM | (LI[\W_]?200)¦(photo.+diode)¦P | 55.0609 | 12.53 |  |  |  | 1 |  | 3% | 2% | 1% | 5% |  | 2.6% | 40 | 1% | 0% | [4], [5], [6], [7]
13 | BNI_A | 0 | 0 | 1 | 2.1237 | 1.597% | 0.02% | 0 | 0% | 0.2% |  | (secondary[\W_]?standard)¦S¦A¦(high[\W_]?quality) |  |  | 13.0488 | 2 | 0.51 | 0.5 |  | 1.5% | 0.1% | 0.2% | 0.5% |  | 0% | 0 | 1% | 0.2% | [1] "High Quality" ~ ISO Secondary Standard
14 | BNI_B | 0 | 0 | 1 | 4.1533 | 1.936% | 0.04% | 0 | 0% | 0.5% |  | ((1¦1st¦I¦first)[\W_]?class)¦B¦(good[\W_]?quality) |  |  | 16.5303 | 4 | 1 | 0.5 |  | 1.5% | 0.5% | 0.5% | 1% |  | 0% | 0 | 2% | 0.5% | [2] "Good Quality" ~ ISO first class
15 | CM22 | 1 | 1 | 0 | 0.5 | 2.352% | 0.01% | 0 | 1% | 0% |  |  | 18.8857 | 4.7312 |  |  |  | 0.5 |  | 2.06% | 0.5% | 0.2% | 1% |  | 1% | 0 | 0.5% | 0% | [7]
16 | CM6b | 1 | 1 | 0 | 0.5 | 2.657% | 0.04% | 0 | 2.236% | 0% |  |  | 23.9593 | 5.3369 |  |  |  | 0.5 |  | 2.09% | 0.5% | 1.2% | 1% |  | 2.236% | 0 | 2% | 0% | [7]
17 | CM3 | 1 | 1 | 0 | 0.5 | 6.423% | 0.12% | 0 | 2.236% | 0% |  |  | 50.7802 | 12.855 |  |  |  | 0.5 |  | 3% | 1% | 2.5% | 5% |  | 2.236% | 0 | 6% | 0% | [7]
18 | TSP800 | 1 | 1 | 0 | 0.5 | 1.897% | 0.02% | 0 | 2.236% | 0% |  | tsp[\W_]?800 | 19.27 | 3.8277 |  |  |  | 0.5 |  | 1.52% | 0.2% | 0.5% | 1% |  | 2.236% | 0 | 1% | 0% | [7]
19 | TSP1 | 1 | 1 | 0 | 0.5 | 3.217% | 0.02% | 0 | 2.236% | 0% |  | tsp[\W_]?1 | 27.6957 | 6.4537 |  |  |  | 0.5 |  | 3.01% | 0.2% | 0.5% | 1% |  | 2.236% | 0 | 1% | 0% | [7]
20 | SPN1 | 1 | 1 | * | 0.5 | 9.597% | 0.02% | 0 | 2.236% | 0% |  |  | 74.5686 | 19.2014 |  |  |  | 0.5 |  | 9.53% | 0.2% | 0.5% | 1% |  | 2.236% | 0 | 1% | 0% | [7]
21 | SPLite | 1 | 1 | 0 | 0.5 | 5.873% | 0.003% | 0 | 2.236% | 0% |  | SP[\W_]?lite | 46.718 | 11.757 |  |  |  | 0.5 |  | 2.12% | 2% | 1% | 5% |  | 2.236% | 0 | 0.15% | 0% | [7]
22 | SP110 | 1 | 1 | 0 | 0.5 | 5.707% | 0.02% | 0 | 2.236% | 0% |  | SP[\W_]?110 | 45.4942 | 11.4249 |  |  |  | 0.5 |  | 2.36% | 1% | 1% | 5% |  | 2.236% | 0 | 1% | 0% | [7]
23 | CH1 | 0 | 0 | 1 | 0.5 | 1.766% | 0.02% | 0 | 2.236% | 0% |  |  |  |  | 22.7997 |  |  | 0.5 |  | 0.78% | 0.1% | 0.5% | 1.5% |  | 2.236% | 0 | 1% | 0% | [7]
24 | NIP2 | 0 | 0 | 1 | 0.5 | 2% | 0.02% | 0 | 1.118% | 0% |  | (NIP[\W#_]?2¦NIP) |  |  | 18.3343 |  |  | 0.5 |  | 1.22% | 0.1% | 0.5% | 1.5% |  | 1.118% | 0 | 1% | 0% | [7]
25 | NIP1 | 0 | 0 | 1 | 0.5 | 1.958% | 0.02% | 0 | 0.5% | 0% |  | NIP[\W#_]?1 |  |  | 16.1719 |  |  | 0.5 |  | 1.15% | 0.1% | 0.5% | 1.5% |  | 0.5% | 0 | 1% | 0% | [7]
26 | LI201 | 1 | 1 | 0 | 2 | 8.84% | 0.02% | 0 | 0.5% | 0% |  | LI[\W_]?201 | 67.7719 | 17.7919 |  |  |  | 2 |  | 3.59% | 0.5% | 1% | 8% |  | 0.5% | 0 | 1% | 0% | [7]
27 | Schenk | 1 | 1 | 0 | 1 | 2.291% | 0.02% | 30 | 3% | 2% |  | (shenk¦schenk) | 29.8956 | 4.6904 |  |  |  | 1 |  | 2% | 1% | 0.5% | 0% |  | 3% | 30 | 1% | 2% | Datasheet, [8]
28 | CHP1 | 0 | 0 | 1 | 1.118 | 0.9644% | 0.01% | 0 | 0.5% | 0% |  |  |  |  | 8.7618 |  | 1 | 0.5 |  | 0.8% | 0.5% | 0.2% | 0% |  | 0.5% | 0 | 0.5% | 0% | Datasheet
29 | CM11 | 1 | 1 | 0 | 14.4568 | 2.934% | 0.06% | 50 | 0% | 0.25% |  |  | 39.2257 | 15.6026 |  | 12 | 8 | 1 |  | 2% | 0.5% | 0.6% | 2% |  | 0% | 50 | 3% | 0.25% | Datasheet
Make sure to extend the list on the “export” page after adding sensors!
Note on AC18: 40 W/m² plus an added 2.5% azimuth factor make a reasonable approx. to Wilbert et al. (2.5 % for 0-75° 4.3-19.6 % for 80-90°), consistent with the plots of Vignola et al.
Note on AC19: 40 W/m² plus an added 2.5% azimuth factor make a reasonable approx. to Wilbert et al. (2.5 % for 0-75° 4.3-19.6 % for 80-90°), consistent with the plots of Vignola et al.

## sources
1 | World Meteorological Organization. 2008. Guide to Meteorological Instruments and Methods of Observation. Geneva, Switzerland: World Meteorological Organization.
2 | Kipp & Zonen CMP/CMA series manual, 2006
3 | A. Driesse, W. Zaaiman, D. Riley, N. Taylor, and J. S. Stein, “Indoor and Outdoor Evaluation of Global Irradiance Sensors,” Hamburg, Germany, Sep. 2015, p. 6.
4 | D. L. King and D. R. Myers, “Silicon-photodiode pyranometers: operational characteristics, historical experiences, and new calibration procedures,” in Conference Record of the Twenty Sixth IEEE Photovoltaic Specialists Conference - 1997, Anaheim, CA, USA, 1997, pp. 1285–1288. doi: 10.1109/PVSC.1997.654323.
5 | F. Vignola, “Removing Systematic Errors from Rotating Shadowband Pyranometer Data.” 2015. Accessed: Aug. 10, 2021. [Online].
6 | S. Wilbert et al., “Uncertainty of rotating shadowband irradiometers and Si-pyranometers including the spectral irradiance error,” Cape Town, South Africa, 2016, p. 150009. doi: 10.1063/1.4949241.
7 | NREL BMS Radiometer Data Uncertainty Analysis Worksheet, 08/20/2014, https://midcdmz.nrel.gov/apps/sitehome.pl?site=BMS
8 | http://solardat.uoregon.edu/Instruments.html